## How to Use
HOW TO USE: CALIBRATION DRIFT AND INTERVAL CALCULATOR
Colour legend
Yellow cells: type your data here.
Grey cells: formulas. They calculate automatically; do not overwrite them.
Steps
1. Enter tolerance, current interval and safety factor.
2. For each calibration enter the date, months since the previous one and the maximum as found error.
3. Read average and worst drift rates.
4. Use the recommended interval as an input to your interval review.
What the formulas do
• Drift rate = ¦as found error¦ ÷ months × 12.
• Recommended = tolerance ÷ worst rate × 12 × safety factor, between 3 and 24 months.
• Any out of tolerance event: recommended = half the current interval (minimum 3).
Sheet "Calculator" holds a worked example: overwrite the yellow cells with your own data.
Basis: Assumes the instrument is adjusted close to zero at each calibration, so the as found error is the drift over the interval. Drift rate = ¦as found error¦ ÷ months × 12. Recommended interval = tolerance ÷ worst drift rate × 12 × safety factor, limited to 3 to 24 months. Engineering judgement required, in the spirit of ILAC G24.
More free templates, calculators and tools: https://instrumentationblog.in/downloads/templates/

## Calculator
CALIBRATION DRIFT AND INTERVAL CALCULATOR
Project: PT 1101 interval review |  |  |  | Calc no: CALC CAL 003 Rev: 0
Yellow cells: enter data. Grey cells: calculated automatically. Free template from instrumentationblog.in
Instrument
Tag |  | PT 1101 |  | Tolerance (± % span) |  | 0.25
Current interval (months) |  | 12 |  | Safety factor (0 to 1) |  | 0.7
Calibration history
No | Calibration date | Months since previous | Max as found error (% span) | Drift rate (% span per year) | Within tolerance | Adjusted to zero? | Remarks
1 | 2022-10-10 | 12 | 0.1 | 0.1 | YES | Yes
2 | 2023-10-12 | 12 | 0.12 | 0.12 | YES | Yes
3 | 2024-10-08 | 12 | 0.09 | 0.09 | YES | Yes
4 | 2025-10-14 | 12 | 0.14 | 0.14 | YES | Yes
5 | 2026-10-12 | 12 | 0.11 | 0.11 | YES | Yes
6
7
8
9
10
Average drift rate (% span per year) | [5 empty] | 0.112
Worst drift rate (% span per year) | [5 empty] | 0.14
Out of tolerance events | [5 empty] | 0
Recommended interval (months) | [5 empty] | 15
If any as found result was out of tolerance, the sheet halves the current interval instead. Review outliers, repairs and changes of service before acting.
Prepared by |  |  |  | Approved by
Name: |  |  |  | Name:
Signature: |  |  |  | Signature:
Date: |  |  |  | Date: