## How to Use
HOW TO USE: MEASUREMENT UNCERTAINTY BUDGET CALCULATOR
Colour legend
Yellow cells: type your data here.
Grey cells: formulas. They calculate automatically; do not overwrite them.
Steps
1. Enter the measurand, unit, coverage factor and the tolerance of the instrument you are calibrating.
2. List every source of uncertainty with its half width (±) and distribution.
3. Enter a sensitivity coefficient if the source is in a different unit; leave blank for 1.
4. Read the combined and expanded uncertainty, and check the TUR.
What the formulas do
• u = value ÷ divisor × sensitivity.
• uc = √(Σ u²), U = k × uc.
• TUR = tolerance ÷ U.
Sheet "Calculator" holds a worked example: overwrite the yellow cells with your own data.
Basis: JCGM 100:2008 (GUM): standard uncertainty from distribution divisors, root sum of squares combination, expanded uncertainty U = k × uc.
More free templates, calculators and tools: https://instrumentationblog.in/downloads/templates/

## Calculator
MEASUREMENT UNCERTAINTY BUDGET CALCULATOR
Project: Digital pressure gauge calibration, 0 to 20 bar |  |  |  | Calc no: CALC UNC 001 Rev: 0
Yellow cells: enter data. Grey cells: calculated automatically. Free template from instrumentationblog.in
Measurement
Measurand |  | Pressure at 10 bar point |  | Unit |  | bar
Coverage factor k |  | 2 |  | Tolerance of instrument under test (±) |  | 0.02
Uncertainty contributions
Source | Value (±) | Distribution | Divisor | Sensitivity ci | Standard uncertainty | u²
Reference standard calibration certificate | 0.004 | Normal (k=2) | 2 | 1 | 0.002 | 4e-06
Reference standard drift since calibration | 0.002 | Rectangular | 1.7321 | 1 | 0.0011547 | 1.3333e-06
Resolution of instrument under test (half digit) | 0.0005 | Rectangular | 1.7321 | 1 | 0.00028868 | 8.3333e-08
Repeatability (standard deviation of readings) | 0.0015 | Normal (k=1) | 1 | 1 | 0.0015 | 2.25e-06
Temperature effect on reference | 0.001 | Rectangular | 1.7321 | 1 | 0.00057735 | 3.3333e-07
Head correction (height difference) | 0.0005 | Rectangular | 1.7321 | 1 | 0.00028868 | 8.3333e-08
Combined standard uncertainty uc | [5 empty] | 0.0028431
Expanded uncertainty U = k × uc | [5 empty] | 0.0056862
Test uncertainty ratio TUR = tolerance ÷ U | [5 empty] | 3.5173
TUR check (≥ 4 recommended) | [5 empty] | CHECK
Divisors: normal k=2 → 2, normal k=1 → 1, rectangular → √3, triangular → √6, U shaped → √2. Values are example figures for illustration.
Prepared by |  |  |  | Checked by
Name: |  |  |  | Name:
Signature: |  |  |  | Signature:
Date: |  |  |  | Date: