| Your Company Name / Logo | CORIOLIS FLOWMETER ZERO CHECK AND VERIFICATION SOP | Doc no: SOP INST 009 | Rev: 0 |
| --- | --- | --- | --- |
| Department: Instrumentation | Standard operating procedure | Date: | Page: see footer |

# 1. Purpose

To confirm that a Coriolis flowmeter reads zero at true no flow and that its sensor tubes are healthy, and to correct the zero only when the check shows it is needed.

# 2. Scope

Coriolis mass flowmeters in liquid and gas service that can be blocked in to a no flow, full condition.

# 3. References

- Flowmeter manufacturer manual (zero and verification sections)

- Instrument data sheet (zero stability, accuracy)

- Site isolation procedure

# 4. Responsibilities

- Instrument technician: carries out the work and records results.

- Instrument engineer or supervisor: reviews and approves the records.

- Operations: grants the permit, isolates the process and returns the equipment to service.

# 5. Safety precautions

- Work only under a valid permit to work and follow the site isolation (lockout tagout) procedure.

- Wear the PPE required for the area: safety helmet, glasses, gloves and safety shoes as a minimum.

- In hazardous areas use certified or intrinsically safe test equipment, or obtain a gas free certificate and hot work permit.

- Inform the control room before starting, and before returning the instrument to service.

- Blocking in a line may affect upstream equipment: agree the isolation plan with operations.

- Do not open the meter housing in a hazardous area unless the permit allows it.

# 6. Tools and equipment

- HART communicator or the meter local display and keypad

- Meter vendor diagnostic software (for verification reports), if used on site

- Stopwatch for zero stability observation

# 7. Prerequisites

- Permit approved; the flow loop and any totaliser based batching put in manual or paused.

- Upstream and downstream block valves available and tight shut off.

- Data sheet zero stability value known (for example in kg/h).

# 8. Procedure

## 8.1 Establish true zero flow

- Run the process at normal temperature and pressure so the meter is at operating conditions.

- Close the downstream block valve first, then the upstream block valve, so the meter stays full and pressurised.

- Confirm there is no flow through bypasses and that the meter is completely full (no gas pockets in liquid service, no condensate in gas service).

- Wait for the temperature to stabilise, typically several minutes.

## 8.2 Zero check (as found)

- Observe the mass flow reading for at least one minute and note the average and the spread.

- Compare the average with the data sheet zero stability. If it is within the zero stability, do not adjust the zero.

- Record the as found zero reading and the density reading for comparison with the expected fluid density.

## 8.3 Zero adjustment (only if required)

- If the as found zero is outside the zero stability and the no flow condition is confirmed, start the zero routine from the communicator or display.

- Do not disturb the meter or the piping while the routine runs.

- Record the new zero value and check the reading again for one minute.

## 8.4 Meter verification

- Run the built in meter verification (tube stiffness or sensor health test) where the meter supports it.

- Record the result (pass or fail) and save or print the verification report.

- If verification fails, report it to the instrument engineer; possible causes are coating, erosion or tube damage.

## 8.5 Return to service

- Open the upstream block valve slowly, then the downstream valve.

- Confirm the flow reading with operations and release the loop to normal operation.

# 9. Acceptance criteria

- As found or as left zero reading within the data sheet zero stability.

- Meter verification result PASS (where available).

- Density reading consistent with the process fluid.

# 10. Records

- Calibration master list (free template: instrumentationblog.in/downloads/templates/calibration-master-list/)

- Instrument PM schedule (free template: instrumentationblog.in/downloads/templates/instrument-preventive-maintenance-schedule/)

# 11. Approval

|  | Prepared by | Reviewed by | Approved by |
| --- | --- | --- | --- |
| Name |  |  |  |
| Signature |  |  |  |
| Date |  |  |  |

This procedure is a generic template. Adapt it to your site permit system, the manufacturer instructions and your quality system before use.