| Your Company Name / Logo | FIXED GAS DETECTOR BUMP TEST AND CALIBRATION SOP | Doc no: SOP ANA 002 | Rev: 0 |
| --- | --- | --- | --- |
| Department: Instrumentation | Standard operating procedure | Date: | Page: see footer |

# 1. Purpose

To make sure fixed gas detectors respond to gas, read correctly and raise their alarms, so leaks are detected in time.

# 2. Scope

Fixed point flammable gas detectors (catalytic and infrared) and toxic gas detectors (electrochemical) with 4 to 20 mA or digital outputs.

# 3. References

- Detector manufacturer manual and calibration adapter instructions

- Test gas certificates

- Fire and gas cause and effect matrix

- Site gas detection philosophy

# 4. Responsibilities

- Instrument technician: carries out the work and records results.

- Instrument engineer or supervisor: reviews and approves the records.

- Operations: grants the permit, isolates the process and returns the equipment to service.

# 5. Safety precautions

- Work only under a valid permit to work and follow the site isolation (lockout tagout) procedure.

- Wear the PPE required for the area: safety helmet, glasses, gloves and safety shoes as a minimum.

- In hazardous areas use certified or intrinsically safe test equipment, or obtain a gas free certificate and hot work permit.

- Inform the control room before starting, and before returning the instrument to service.

- Inhibit only the detector under test and keep the inhibit time short; the area loses protection while inhibited.

- Handle test gas cylinders upright and secured; toxic test gas must be vented away from people.

# 6. Tools and equipment

- Certified test gas of the correct gas and concentration (typically about 50 % LEL for flammable, near the alarm level for toxic) within expiry

- Zero gas (synthetic air or nitrogen as required) or clean air

- Calibration adapter (cup) and regulator with the flow rate given by the manufacturer

- Magnet or communicator to start calibration, and a radio

# 7. Prerequisites

- Permit approved; fire and gas system detector inhibited by operations.

- Test gas certificate checked; the detector sensor is not past its replacement date.

# 8. Procedure

## 8.1 Bump test (as found)

- Fit the calibration adapter and apply test gas at the manufacturer flow rate.

- Record the peak reading and the time to reach it; compare with the gas concentration.

- If the reading is within the acceptance band the detector passes the bump test; continue to calibration only if required by the schedule or if the bump test fails.

## 8.2 Zero calibration

- Apply zero gas or ensure clean air, start the zero routine and wait for it to finish.

- Record the zero reading.

## 8.3 Span calibration

- Enter the test gas concentration if required, apply test gas and start the span routine.

- Wait for the detector to accept the span, then remove the gas and record the as left span reading.

- Allow the reading to return to zero before ending the inhibit.

## 8.4 Alarm and output check

- With the control room, confirm the detector signal (mA or digital value) matches the local reading during gas application.

- Confirm the low and high alarms were seen at their set points where the test gas allows, or simulate the signal to check alarm actions.

- Remove the inhibit and confirm the detector is healthy in the system.

# 9. Acceptance criteria

- Bump test reading within the site acceptance band (for example within ± 10 % of the applied concentration, or as set by the site procedure).

- As left span reading within the manufacturer tolerance.

- Zero stable at zero, and the control system reading matches the detector.

# 10. Records

- Calibration master list (free template: instrumentationblog.in/downloads/templates/calibration-master-list/)

- Instrument PM schedule (free template: instrumentationblog.in/downloads/templates/instrument-preventive-maintenance-schedule/)

# 11. Approval

|  | Prepared by | Reviewed by | Approved by |
| --- | --- | --- | --- |
| Name |  |  |  |
| Signature |  |  |  |
| Date |  |  |  |

This procedure is a generic template. Adapt it to your site permit system, the manufacturer instructions and your quality system before use.