| Your Company Name / Logo | WEIGHING SYSTEM CALIBRATION SOP | Doc no: SOP INST 019 | Rev: 0 |
| --- | --- | --- | --- |
| Department: Instrumentation | Standard operating procedure | Date: | Page: see footer |

# 1. Purpose

To make weighing systems read mass correctly for batching, inventory and custody use.

# 2. Scope

Tank, hopper and platform weighing systems with strain gauge load cells.

# 3. References

- Weighing system manual

- Test weight certificates

- Mechanical drawings of vessel supports

# 4. Responsibilities

- Instrument technician: carries out the work and records results.

- Instrument engineer or supervisor: reviews and approves the records.

- Operations: grants the permit, isolates the process and returns the equipment to service.

# 5. Safety precautions

- Work only under a valid permit to work and follow the site isolation (lockout tagout) procedure.

- Wear the PPE required for the area: safety helmet, glasses, gloves and safety shoes as a minimum.

- In hazardous areas use certified or intrinsically safe test equipment, or obtain a gas free certificate and hot work permit.

- Inform the control room before starting, and before returning the instrument to service.

- Use certified lifting gear for test weights; keep clear of suspended loads.

- Isolate filling and discharge to prevent material movement.

# 6. Tools and equipment

- Certified test weights or a calibrated flow meter for substitution

- Weighing indicator or configuration software

- Spirit level and torch

# 7. Prerequisites

- Vessel empty or at known contents; process isolated.

# 8. Procedure

## 8.1 Mechanical checks

- Check flexible connections, no bridging, mounts level and free, stays loose as designed.

- Check cables and junction box.

## 8.2 Zero

- With the empty vessel, set or record zero.

## 8.3 Load test

- Apply known loads in steps up to near capacity, recording readings.

- Remove loads in the same steps.

- If weights are not enough, use build up with substitution.

## 8.4 Adjust and verify

- Adjust span if needed and repeat.

- Record results and return to service.

# 9. Acceptance criteria

- Error and hysteresis within tolerance (for example ± 0.1 % of capacity).

- Mechanical defects corrected and recorded.

# 10. Records

- Load cell calibration record (free template: instrumentationblog.in/downloads/templates/load-cell-calibration-record/)

- Calibration master list (free template: instrumentationblog.in/downloads/templates/calibration-master-list/)

# 11. Approval

|  | Prepared by | Reviewed by | Approved by |
| --- | --- | --- | --- |
| Name |  |  |  |
| Signature |  |  |  |
| Date |  |  |  |

This procedure is a generic template. Adapt it to your site permit system, the manufacturer instructions and your quality system before use.