| Your Company Name / Logo | MAGNETIC FLOWMETER VERIFICATION SOP | Doc no: SOP INST 008 | Rev: 0 |
| --- | --- | --- | --- |
| Department: Instrumentation | Standard operating procedure | Date: | Page: see footer |

# 1. Purpose

To confirm that a magnetic flowmeter is measuring correctly without removing it for a wet calibration.

# 2. Scope

Electromagnetic flowmeters in conductive liquid service with 4 to 20 mA, pulse or digital outputs.

# 3. References

- Flowmeter manual and the manufacturer's verification procedure

- Instrument data sheet

# 4. Responsibilities

- Instrument technician: carries out the work and records results.

- Instrument engineer or supervisor: reviews and approves the records.

- Operations: grants the permit, isolates the process and returns the equipment to service.

# 5. Safety precautions

- Work only under a valid permit to work and follow the site isolation (lockout tagout) procedure.

- Wear the PPE required for the area: safety helmet, glasses, gloves and safety shoes as a minimum.

- In hazardous areas use certified or intrinsically safe test equipment, or obtain a gas free certificate and hot work permit.

- Inform the control room before starting, and before returning the instrument to service.

- Opening the sensor or transmitter terminal box may require power isolation and a gas free certificate in hazardous areas.

# 6. Tools and equipment

- Manufacturer verification tool or software (where available)

- Multimeter and insulation tester (only as allowed by the manual)

- mA meter or loop calibrator

# 7. Prerequisites

- Process able to provide a full pipe with zero flow for the zero check (both valves closed).

# 8. Procedure

## 8.1 Visual and installation checks

- Check that the pipe is full at the meter (no air), grounding rings or electrodes are connected, and cables are intact.

## 8.2 Zero check

- Stop the flow with the pipe full (close downstream valve, then upstream).

- Check the reading is close to zero; perform a zero adjustment only if the manual recommends it.

## 8.3 Verification

- Run the built in verification (for example coil resistance, electrode integrity, transmitter electronics).

- Save or print the verification report.

- If no built in tool exists, measure coil resistance and electrode resistance as described in the manual and compare with the factory values.

## 8.4 Output check

- Simulate a flow value from the transmitter and check the mA output and the DCS reading at 0, 50 and 100 %.

# 9. Acceptance criteria

- Verification passed (or measured values within the manual limits).

- Zero reading within the manufacturer specification.

- Output and DCS reading within loop tolerance.

# 10. Records

- Calibration master list (free template: instrumentationblog.in/downloads/templates/calibration-master-list/)

- Instrument breakdown report (if faults found) (free template: instrumentationblog.in/downloads/templates/instrument-breakdown-report/)

# 11. Approval

|  | Prepared by | Reviewed by | Approved by |
| --- | --- | --- | --- |
| Name |  |  |  |
| Signature |  |  |  |
| Date |  |  |  |

This procedure is a generic template. Adapt it to your site permit system, the manufacturer instructions and your quality system before use.