| Your Company Name / Logo | PRESSURE GAUGE CALIBRATION SOP | Doc no: SOP INST 005 | Rev: 0 |
| --- | --- | --- | --- |
| Department: Instrumentation | Standard operating procedure | Date: | Page: see footer |

# 1. Purpose

To confirm that pressure gauges indicate within their accuracy class.

# 2. Scope

Dial pressure and vacuum gauges (EN 837-1 classes 0.6 to 2.5) calibrated in the workshop.

# 3. References

- EN 837-1 (Bourdon tube pressure gauges)

- Comparator or dead weight tester manual

# 4. Responsibilities

- Instrument technician: carries out the work and records results.

- Instrument engineer or supervisor: reviews and approves the records.

- Operations: grants the permit, isolates the process and returns the equipment to service.

# 5. Safety precautions

- Work only under a valid permit to work and follow the site isolation (lockout tagout) procedure.

- Wear the PPE required for the area: safety helmet, glasses, gloves and safety shoes as a minimum.

- In hazardous areas use certified or intrinsically safe test equipment, or obtain a gas free certificate and hot work permit.

- Inform the control room before starting, and before returning the instrument to service.

- Use the correct test medium: oil free (water or air) for oxygen service gauges, never oil.

- Do not exceed the full scale of the gauge by more than its stated overpressure rating.

# 6. Tools and equipment

- Hydraulic or pneumatic comparator, or dead weight tester

- Reference gauge at least 4 times more accurate

- Adapters, seals and tools for the gauge connection

# 7. Prerequisites

- Gauge removed under permit, cleaned of process fluid.

- Gauge visually checked: glass, pointer, dial and connection undamaged.

# 8. Procedure

## 8.1 Setup

- Mount the gauge vertically (its normal position) next to the reference.

- Exercise the gauge to full scale and back to zero three times.

## 8.2 Test

- Check the pointer at zero with no pressure applied.

- Apply 0, 25, 50, 75 and 100 % of full scale upscale, tapping the gauge lightly before reading.

- Return downscale through the same points without overshooting.

- Record readings and calculate error and hysteresis in % of span.

## 8.3 Adjustment

- If zero is out, reset the pointer (adjustable pointer or micrometer pointer).

- If span or linearity are out, adjust the movement if the design allows; otherwise replace the gauge.

- Repeat the test after adjustment.

# 9. Acceptance criteria

- Error and hysteresis within the accuracy class at every point.

- Pointer returns to zero.

# 10. Records

- Pressure gauge calibration report (free template: instrumentationblog.in/downloads/templates/pressure-gauge-calibration-report/)

# 11. Approval

|  | Prepared by | Reviewed by | Approved by |
| --- | --- | --- | --- |
| Name |  |  |  |
| Signature |  |  |  |
| Date |  |  |  |

This procedure is a generic template. Adapt it to your site permit system, the manufacturer instructions and your quality system before use.