| Your Company Name / Logo | TEMPERATURE ELEMENT REPLACEMENT SOP | Doc no: SOP INST 017 | Rev: 0 |
| --- | --- | --- | --- |
| Department: Instrumentation | Standard operating procedure | Date: | Page: see footer |

# 1. Purpose

To replace failed temperature elements without opening the process and with a correct, verified measurement afterwards.

# 2. Scope

RTD and thermocouple insert elements in thermowells, with head or remote transmitters.

# 3. References

- Temperature transmitter and thermowell data sheets

- Spare parts list

- Loop diagram

# 4. Responsibilities

- Instrument technician: carries out the work and records results.

- Instrument engineer or supervisor: reviews and approves the records.

- Operations: grants the permit, isolates the process and returns the equipment to service.

# 5. Safety precautions

- Work only under a valid permit to work and follow the site isolation (lockout tagout) procedure.

- Wear the PPE required for the area: safety helmet, glasses, gloves and safety shoes as a minimum.

- In hazardous areas use certified or intrinsically safe test equipment, or obtain a gas free certificate and hot work permit.

- Inform the control room before starting, and before returning the instrument to service.

- Never remove a thermowell in service; only the element inside the well.

- Elements from hot wells are hot: use gloves.

- If the well is suspected leaking, stop and treat it as a pressure boundary failure.

# 6. Tools and equipment

- Replacement element of correct type, length and diameter

- Multimeter or temperature calibrator

- Spanners and heat resistant gloves

# 7. Prerequisites

- Loop in manual; alarms and trips inhibited as authorised.

# 8. Procedure

## 8.1 Remove

- Disconnect wires, noting terminal positions.

- Unscrew and withdraw the element; check the well bore for liquid, which indicates a leaking well.

## 8.2 Check new element

- Confirm type (Pt100, type K), wiring, length and diameter.

- Measure resistance or mV at ambient and compare with the table.

## 8.3 Install

- Insert until the tip touches the well bottom (spring loaded).

- Reconnect wires with correct polarity or 3 wire arrangement.

## 8.4 Verify

- Compare the reading with a nearby instrument or a portable probe.

- Remove inhibits and record the replacement.

# 9. Acceptance criteria

- Correct element type and length installed, touching the well bottom.

- Reading consistent with process and nearby instruments.

# 10. Records

- Instrument maintenance log (free template: instrumentationblog.in/downloads/templates/instrument-maintenance-log/)

- Instrument spare parts list (free template: instrumentationblog.in/downloads/templates/instrument-spare-parts-list/)

# 11. Approval

|  | Prepared by | Reviewed by | Approved by |
| --- | --- | --- | --- |
| Name |  |  |  |
| Signature |  |  |  |
| Date |  |  |  |

This procedure is a generic template. Adapt it to your site permit system, the manufacturer instructions and your quality system before use.