| Your Company Name / Logo | RTD AND THERMOCOUPLE SENSOR CALIBRATION SOP | Doc no: SOP INST 004 | Rev: 0 |
| --- | --- | --- | --- |
| Department: Instrumentation | Standard operating procedure | Date: | Page: see footer |

# 1. Purpose

To verify that temperature sensors are within their tolerance class and to detect drift or damage.

# 2. Scope

Pt100 RTDs and base metal thermocouples (K, J, T, N, E) removed from their thermowells, tested from ambient to the dry block limit.

# 3. References

- IEC 60751 (RTD tolerance classes)

- IEC 60584-1 (thermocouple tolerances)

- Dry block calibrator manual

# 4. Responsibilities

- Instrument technician: carries out the work and records results.

- Instrument engineer or supervisor: reviews and approves the records.

- Operations: grants the permit, isolates the process and returns the equipment to service.

# 5. Safety precautions

- Work only under a valid permit to work and follow the site isolation (lockout tagout) procedure.

- Wear the PPE required for the area: safety helmet, glasses, gloves and safety shoes as a minimum.

- In hazardous areas use certified or intrinsically safe test equipment, or obtain a gas free certificate and hot work permit.

- Inform the control room before starting, and before returning the instrument to service.

- Dry blocks reach several hundred °C: use heat resistant gloves, let probes cool before handling, and keep the block on a stable, clear surface.

# 6. Tools and equipment

- Dry block calibrator with suitable insert, or stirred liquid bath

- Reference probe and readout with valid certificate

- Resistance or mV readout for the sensor under test (or the loop transmitter)

- Heat resistant gloves

# 7. Prerequisites

- Sensor removed from the thermowell under permit; process side safely isolated.

- Insert bore matching the sensor diameter.

# 8. Procedure

## 8.1 Setup

- Insert the reference probe and the sensor under test into adjacent bores to the full depth of the insert.

- Make sure the immersion is at least 15 times the sheath diameter where possible, to avoid stem conduction error.

## 8.2 Test

- Choose 3 points covering the operating range (for example 0 or ambient, mid range, and near maximum).

- At each point wait for stability (typically 10 to 20 minutes, or as the dry block indicates stable).

- Record the reference temperature and the sensor reading (Ω or mV converted to °C).

- Calculate the error and compare with the tolerance class for that temperature.

## 8.3 Completion

- Allow the sensor to cool before handling and reinstalling.

- Reinstall in the thermowell with the correct immersion, and check the reading in the DCS.

# 9. Acceptance criteria

- RTD error within its class (for example class A: ± (0.15 + 0.002 |t|) °C; class B: ± (0.3 + 0.005 |t|) °C).

- Thermocouple error within its IEC 60584-1 class tolerance.

- Sensor insulation and continuity normal.

# 10. Records

- Calibration certificate format (free template: instrumentationblog.in/downloads/templates/calibration-certificate-format/)

- Calibration master list (free template: instrumentationblog.in/downloads/templates/calibration-master-list/)

# 11. Approval

|  | Prepared by | Reviewed by | Approved by |
| --- | --- | --- | --- |
| Name |  |  |  |
| Signature |  |  |  |
| Date |  |  |  |

This procedure is a generic template. Adapt it to your site permit system, the manufacturer instructions and your quality system before use.