| Your Company Name / Logo | TEMPERATURE TRANSMITTER CALIBRATION SOP | Doc no: SOP INST 003 | Rev: 0 |
| --- | --- | --- | --- |
| Department: Instrumentation | Standard operating procedure | Date: | Page: see footer |

# 1. Purpose

To verify that temperature transmitters convert the sensor signal to the correct 4 to 20 mA output within tolerance.

# 2. Scope

Temperature transmitters with Pt100 (2, 3 or 4 wire) or thermocouple inputs and 4 to 20 mA or HART output. Sensor checks are covered in a separate SOP.

# 3. References

- IEC 60751 (Pt100 tables)

- IEC 60584-1 (thermocouple tables)

- Transmitter manual and data sheet

# 4. Responsibilities

- Instrument technician: carries out the work and records results.

- Instrument engineer or supervisor: reviews and approves the records.

- Operations: grants the permit, isolates the process and returns the equipment to service.

# 5. Safety precautions

- Work only under a valid permit to work and follow the site isolation (lockout tagout) procedure.

- Wear the PPE required for the area: safety helmet, glasses, gloves and safety shoes as a minimum.

- In hazardous areas use certified or intrinsically safe test equipment, or obtain a gas free certificate and hot work permit.

- Inform the control room before starting, and before returning the instrument to service.

- Open transmitter heads in hazardous areas only when the area is gas free or the circuit is intrinsically safe and the permit allows it.

# 6. Tools and equipment

- RTD and thermocouple simulator / calibrator with cold junction compensation

- mA meter or documenting calibrator

- HART communicator

- Correct thermocouple extension or compensating wire for the simulator connection

# 7. Prerequisites

- Loop in manual and trips bypassed as approved.

- Sensor type, wiring and range confirmed from the data sheet.

# 8. Procedure

## 8.1 Preparation

- Disconnect the sensor wires from the transmitter and label them.

- Connect the calibrator to the input terminals using the same wiring (3 wire for a 3 wire RTD; thermocouple wire of the same type for thermocouples).

- For thermocouples, use the calibrator's internal cold junction compensation, or simulate the mV value directly with the transmitter's CJC disabled as per the manual.

## 8.2 As found test

- Simulate 0, 25, 50, 75 and 100 % of the calibrated range and record the output.

- Record the simulated value (Ω or mV) and the expected mA.

- Calculate errors in % of span.

## 8.3 Adjustment and as left

- If out of tolerance, perform the input (sensor) trim and/or output trim as described in the manual.

- Repeat the test and record the as left values.

## 8.4 Return to service

- Reconnect the sensor wires exactly as before; check terminal tightness.

- Close the head cover with its seal in place.

- Confirm a sensible temperature reading in the DCS.

# 9. Acceptance criteria

- As left errors within tolerance (typically ± 0.25 % of span).

- Sensor reconnected and reading correct.

# 10. Records

- Temperature transmitter calibration report (free template: instrumentationblog.in/downloads/templates/temperature-transmitter-calibration-report/)

# 11. Approval

|  | Prepared by | Reviewed by | Approved by |
| --- | --- | --- | --- |
| Name |  |  |  |
| Signature |  |  |  |
| Date |  |  |  |

This procedure is a generic template. Adapt it to your site permit system, the manufacturer instructions and your quality system before use.