| Your Company Name / Logo | VORTEX FLOWMETER VERIFICATION SOP | Doc no: SOP INST 013 | Rev: 0 |
| --- | --- | --- | --- |
| Department: Instrumentation | Standard operating procedure | Date: | Page: see footer |

# 1. Purpose

To confirm a vortex meter is configured correctly and measuring reliably, without removing it from the line.

# 2. Scope

Inline vortex flowmeters on steam, gas and liquid service with 4 to 20 mA, pulse or HART outputs.

# 3. References

- Meter manual and calibration certificate (K factor)

- Flow data sheet and design conditions

- Loop diagram

# 4. Responsibilities

- Instrument technician: carries out the work and records results.

- Instrument engineer or supervisor: reviews and approves the records.

- Operations: grants the permit, isolates the process and returns the equipment to service.

# 5. Safety precautions

- Work only under a valid permit to work and follow the site isolation (lockout tagout) procedure.

- Wear the PPE required for the area: safety helmet, glasses, gloves and safety shoes as a minimum.

- In hazardous areas use certified or intrinsically safe test equipment, or obtain a gas free certificate and hot work permit.

- Inform the control room before starting, and before returning the instrument to service.

- Steam lines are hot; do not open the meter body in service.

# 6. Tools and equipment

- HART communicator or configuration software

- mA calibrator or multimeter

- Frequency source if electronics simulation is used

# 7. Prerequisites

- Operations aware; flow loop may be put in manual if simulation is used.

# 8. Procedure

## 8.1 Configuration check

- Compare K factor, pipe ID, fluid, density basis and units with the calibration certificate and data sheet.

- Check damping, low flow cutoff and output range.

## 8.2 Electronics test

- Run the built in self test or simulate a shedding frequency.

- Confirm the displayed flow and output current match the expected values.

## 8.3 Signal quality

- Review signal strength or noise diagnostics at normal flow.

- Look for vibration sources and check the meter does not read flow when the line is stopped.

## 8.4 Process comparison

- Compare the reading with mass balance, a check meter or other plant data.

- Record findings and restore the loop to automatic.

# 9. Acceptance criteria

- Configuration matches the certificate and data sheet.

- Simulated flow and output within electronics tolerance.

- No false flow at zero; reading consistent with process data.

# 10. Records

- Flowmeter verification record (free template: instrumentationblog.in/downloads/templates/flowmeter-verification-record/)

- Instrument maintenance log (free template: instrumentationblog.in/downloads/templates/instrument-maintenance-log/)

# 11. Approval

|  | Prepared by | Reviewed by | Approved by |
| --- | --- | --- | --- |
| Name |  |  |  |
| Signature |  |  |  |
| Date |  |  |  |

This procedure is a generic template. Adapt it to your site permit system, the manufacturer instructions and your quality system before use.