## How to Use
HOW TO USE: PH ANALYSER CALIBRATION RECORD
Colour legend
Yellow cells: type your data here.
Grey cells: formulas. They calculate automatically; do not overwrite them.
Steps
1. Clean and rinse the electrode.
2. Place it in buffer 1 (pH 7), wait for stability and record mV; repeat with buffer 2.
3. The sheet calculates slope %, offset and electrode condition.
4. Calibrate the analyser, then verify in buffers and record readings.
5. Replace the electrode if slope or offset is outside limits after cleaning.
What the formulas do
• Slope = (mV₂ − mV₁) ÷ (pH₁ − pH₂).
• Theoretical slope = 0.198416 × (T + 273.15) mV/pH.
• Offset at pH 7 = mV₁ + (pH₁ − 7) × slope.
Sheets: "Template" is blank and ready to use. "Example" shows a completed form so you can see the expected format.
More free templates, calculators and tools: https://instrumentationblog.in/downloads/templates/

## Template
PH ANALYSER CALIBRATION RECORD
Company: Your Company Name / Logo |  |  |  | Doc no: Rev:
Yellow cells: enter data. Grey cells: calculated automatically. Free template from instrumentationblog.in
Analyser and calibration
Tag number |  |  |  | Service
Buffer temperature (°C) |  |  |  | Date
Buffer 1 pH |  |  |  | Buffer 1 reading (mV)
Buffer 2 pH |  |  |  | Buffer 2 reading (mV)
Measured slope (mV/pH) |  |  |  | Slope (% of theoretical)
Offset at pH 7 (mV) |  |  |  | Electrode condition
Buffer verification after calibration
No | Buffer pH | Analyser reading (pH) | Error (pH) | Tolerance (± pH) | Result | Buffer lot / expiry | Remarks
1
2
3
4
Verification failures | [5 empty] | 0
Typical acceptance: slope about 90 to 105 % and offset within about ± 30 mV; use your site procedure. Rinse the electrode between buffers and use fresh buffers.
Performed by |  | Checked by |  | Approved by
Name: |  | Name: |  | Name:
Signature: |  | Signature: |  | Signature:
Date: |  | Date: |  | Date:

## Example
PH ANALYSER CALIBRATION RECORD
Company: ABC Process Industries Ltd |  |  |  | Doc no: CAL AT 001 Rev: 0
Yellow cells: enter data. Grey cells: calculated automatically. Free template from instrumentationblog.in
Analyser and calibration
Tag number |  | AT 5101 |  | Service |  | Effluent pH
Buffer temperature (°C) |  | 25 |  | Date |  | 2026-10-25
Buffer 1 pH |  | 7 |  | Buffer 1 reading (mV) |  | -4
Buffer 2 pH |  | 4.01 |  | Buffer 2 reading (mV) |  | 165
Measured slope (mV/pH) |  | 56.5217 |  | Slope (% of theoretical) |  | 95.5441
Offset at pH 7 (mV) |  | -4 |  | Electrode condition |  | GOOD
Buffer verification after calibration
No | Buffer pH | Analyser reading (pH) | Error (pH) | Tolerance (± pH) | Result | Buffer lot / expiry | Remarks
1 | 7 | 7.01 | 0.01 | 0.05 | PASS | L2207 / 2027 03
2 | 4.01 | 4.03 | 0.02 | 0.05 | PASS | L2194 / 2027 01
3 | 10.01 | 9.97 | -0.04 | 0.1 | PASS | L2210 / 2026 12
4
Verification failures | [5 empty] | 0
Typical acceptance: slope about 90 to 105 % and offset within about ± 30 mV; use your site procedure. Rinse the electrode between buffers and use fresh buffers.
Performed by |  | Checked by |  | Approved by
Name: |  | Name: |  | Name:
Signature: |  | Signature: |  | Signature:
Date: |  | Date: |  | Date: