## How to Use
HOW TO USE: PRESSURE GAUGE CALIBRATION REPORT
Colour legend
Yellow cells: type your data here.
Grey cells: formulas. They calculate automatically; do not overwrite them.
Steps
1. Enter the gauge range, unit and accuracy class (for example 1.0 for class 1.0 per EN 837-1, which allows ± 1 % of span).
2. Mount the gauge vertically on the comparator or dead weight tester next to the reference standard.
3. Exercise the gauge to full scale three times, then record readings upscale at each test point.
4. Go to full scale and come back down, recording the downscale readings.
5. Review error, hysteresis and the result, then sign the report. Replace or repair a gauge that fails.
What the formulas do
• Standard pressure = minimum + span × test point % ÷ 100.
• Error (% span) = (gauge reading − standard) ÷ span × 100.
• Hysteresis = ¦upscale − downscale¦ ÷ span × 100.
• PASS when errors and hysteresis are within the accuracy class.
Sheets: "Template" is blank and ready to use. "Example" shows a completed form so you can see the expected format.
More free templates, calculators and tools: https://instrumentationblog.in/downloads/templates/

## Template
PRESSURE GAUGE CALIBRATION REPORT
Company: Your Company Name / Logo |  |  |  | Doc no: Rev:
Yellow cells: enter data. Grey cells: calculated automatically. Free template from instrumentationblog.in
Gauge details
Tag number |  |  |  | Service
Make |  |  |  | Serial number
Range minimum |  |  |  | Range maximum
Unit |  |  |  | Accuracy class (± % of span)
Dial size (mm) |  |  |  | Connection
Test equipment
Reference standard |  |  |  | Standard certificate no
Calibration date |  |  |  | Test medium
Calibration results
Test point (% of span) | Standard pressure | Gauge reading upscale | Error upscale (% span) | Gauge reading downscale | Error downscale (% span) | Hysteresis (% span) | Result
0
25
50
75
100
Maximum error (% span)
Maximum hysteresis (% span)
Overall result
Tap the gauge lightly before each reading. Approach every point from below on the upscale run and from above on the downscale run without overshooting.
Performed by |  | Checked by |  | Approved by
Name: |  | Name: |  | Name:
Signature: |  | Signature: |  | Signature:
Date: |  | Date: |  | Date:

## Example
PRESSURE GAUGE CALIBRATION REPORT
Company: ABC Process Industries Ltd |  |  |  | Doc no: CAL PI 1203 01 Rev: 0
Yellow cells: enter data. Grey cells: calculated automatically. Free template from instrumentationblog.in
Gauge details
Tag number |  | PI 1203 |  | Service |  | Cooling water supply
Make |  | WIKA 232.50 |  | Serial number |  | W23 77810
Range minimum |  | 0 |  | Range maximum |  | 10
Unit |  | bar |  | Accuracy class (± % of span) |  | 1
Dial size (mm) |  | 100 |  | Connection |  | 1/2" NPT bottom
Test equipment
Reference standard |  | Digital test gauge, 0 to 20 bar |  | Standard certificate no |  | NABL/CC 24311
Calibration date |  | 2026-10-12 |  | Test medium |  | Hydraulic oil
Calibration results
Test point (% of span) | Standard pressure | Gauge reading upscale | Error upscale (% span) | Gauge reading downscale | Error downscale (% span) | Hysteresis (% span) | Result
0 | 0 | 0 | 0 | 0.05 | 0.5 | 0.5 | PASS
25 | 2.5 | 2.55 | 0.5 | 2.58 | 0.8 | 0.3 | PASS
50 | 5 | 5.05 | 0.5 | 5.08 | 0.8 | 0.3 | PASS
75 | 7.5 | 7.55 | 0.5 | 7.58 | 0.8 | 0.3 | PASS
100 | 10 | 10.05 | 0.5 | 10.05 | 0.5 | 0 | PASS
Maximum error (% span) | [5 empty] | 0.8
Maximum hysteresis (% span) | [5 empty] | 0.5
Overall result | [5 empty] | PASS
Tap the gauge lightly before each reading. Approach every point from below on the upscale run and from above on the downscale run without overshooting.
Performed by |  | Checked by |  | Approved by
Name: |  | Name: |  | Name:
Signature: |  | Signature: |  | Signature:
Date: |  | Date: |  | Date: