## How to Use
HOW TO USE: PRESSURE TRANSMITTER DATA SHEET
Colour legend
Yellow cells: type your data here.
Grey cells: formulas. They calculate automatically; do not overwrite them.
Steps
1. Fill the general and process sections from the process data sheet and P&ID; always give normal, maximum and design values.
2. Choose a measuring range so the calibrated range sits comfortably inside it, typically 1:1 to 10:1 turndown.
3. Specify the Ex certification to match the hazardous area classification.
4. Select wetted materials for the process fluid (SS 316L for most services, Hastelloy C or tantalum for corrosive fluids).
5. Issue the data sheet with the requisition; the vendor completes the make and model.
What the formulas do
• No formulas: specification document. Section rows organise the data like an ISA 20 style data sheet.
Sheets: "Template" is blank and ready to use. "Example" shows a completed form so you can see the expected format.
More free templates, calculators and tools: https://instrumentationblog.in/downloads/templates/

## Template
PRESSURE TRANSMITTER DATA SHEET
Company: Your Company Name / Logo |  |  | Doc no: Rev:
Yellow cells: enter data. Grey cells: calculated automatically. Free template from instrumentationblog.in
Document
Project |  |  | Revision
Date |  |  | Requisition no
Specification
No | Parameter |  | Value | Unit / notes
General
1 | Tag number
2 | Service
3 | Line / equipment number
4 | P&ID number
5 | Hazardous area classification
Process conditions
6 | Fluid
7 | Operating pressure (normal / max)
8 | Design pressure
9 | Operating temperature (normal / max)
10 | Design temperature
Transmitter
11 | Type
12 | Measuring range (sensor limits)
13 | Calibrated range
14 | Output and protocol
15 | Accuracy
16 | Power supply
17 | Enclosure and ingress protection
18 | Ex certification
19 | Electrical entry
20 | Process connection
21 | Wetted parts / diaphragm material
22 | Fill fluid
23 | Display
Accessories
24 | Manifold
25 | Mounting bracket
26 | Diaphragm seal
Purchase
27 | Make and model
28 | Quantity
Prepared by |  | Checked by | Approved by
Name: |  | Name: | Name:
Signature: |  | Signature: | Signature:
Date: |  | Date: | Date:

## Example
PRESSURE TRANSMITTER DATA SHEET
Company: ABC Process Industries Ltd |  |  | Doc no: DS PT 01 Rev: 0
Yellow cells: enter data. Grey cells: calculated automatically. Free template from instrumentationblog.in
Document
Project |  | Boiler 3 upgrade | Revision | B
Date |  | 2026-10-12 | Requisition no | REQ INST 012
Specification
No | Parameter |  | Value |  | Unit / notes
General
1 | Tag number |  | PT 1101
2 | Service |  | Steam header pressure
3 | Line / equipment number |  | 8" ST 1101
4 | P&ID number |  | PID 100 02
5 | Hazardous area classification |  | Zone 2, IIB T3 |  | IEC 60079-10-1
Process conditions
6 | Fluid |  | Saturated steam
7 | Operating pressure (normal / max) |  | 10 / 12 |  | bar g
8 | Design pressure |  | 16 |  | bar g
9 | Operating temperature (normal / max) |  | 184 / 192 |  | °C
10 | Design temperature |  | 220 |  | °C
Transmitter
11 | Type |  | Gauge pressure, smart
12 | Measuring range (sensor limits) |  | 0 to 55 |  | bar g
13 | Calibrated range |  | 0 to 16 |  | bar g
14 | Output and protocol |  | 4 to 20 mA HART 7
15 | Accuracy |  | ± 0.075 % of span |  | Reference conditions
16 | Power supply |  | 10.5 to 42.4 |  | V DC
17 | Enclosure and ingress protection |  | Aluminium, IP66/67
18 | Ex certification |  | Ex db IIC T6 Gb |  | ATEX / IECEx
19 | Electrical entry |  | M20 x 1.5
20 | Process connection |  | 1/2" NPT female
21 | Wetted parts / diaphragm material |  | SS 316L
22 | Fill fluid |  | Silicone oil
23 | Display |  | LCD, local
Accessories
24 | Manifold |  | 2 valve, SS 316
25 | Mounting bracket |  | 2" pipe mount, SS
26 | Diaphragm seal |  | Not required |  | Use siphon (pigtail) for steam
Purchase
27 | Make and model |  | Vendor to specify
28 | Quantity |  | 1 |  | no
Prepared by |  | Checked by | Approved by
Name: |  | Name: | Name:
Signature: |  | Signature: | Signature:
Date: |  | Date: | Date: