## How to Use
HOW TO USE: RTD AND THERMOCOUPLE SENSOR CALIBRATION REPORT
Colour legend
Yellow cells: type your data here.
Grey cells: formulas. They calculate automatically; do not overwrite them.
Steps
1. Enter the sensor type and tolerance, for example IEC 60751 class A at the highest test temperature.
2. Insert sensor and reference in the dry block at the same depth.
3. At each set point wait for stability, then record reference and sensor readings.
4. Record the raw Ω or mV value to help diagnose faults.
5. Review errors and decide: accept, correct in transmitter, or replace.
What the formulas do
• Error (°C) = sensor reading − reference reading.
• PASS when ¦error¦ ≤ tolerance.
Sheets: "Template" is blank and ready to use. "Example" shows a completed form so you can see the expected format.
More free templates, calculators and tools: https://instrumentationblog.in/downloads/templates/

## Template
RTD AND THERMOCOUPLE SENSOR CALIBRATION REPORT
Company: Your Company Name / Logo |  |  |  | Doc no: Rev:
Yellow cells: enter data. Grey cells: calculated automatically. Free template from instrumentationblog.in
Sensor details
Tag number |  |  |  | Sensor type
Use range |  |  |  | Tolerance (± °C)
Temperature source |  |  |  | Reference thermometer
Calibration date |  |  |  | Immersion depth (mm)
Calibration results
Point | Set temperature (°C) | Reference (°C) | Sensor as found (°C) | As found error (°C) | As found result | Raw value (Ω or mV) | Stable (yes/no) | Remarks
1
2
3
4
5
6
Maximum error (°C)
Overall result
Wait for the block and both thermometers to stabilise at each point. Use an insert that fits the sensor closely and immerse it fully. Sensors cannot be adjusted; a failed sensor is replaced or corrected in the transmitter.
Performed by |  |  | Checked by |  |  | Approved by
Name: |  |  | Name: |  |  | Name:
Signature: |  |  | Signature: |  |  | Signature:
Date: |  |  | Date: |  |  | Date:

## Example
RTD AND THERMOCOUPLE SENSOR CALIBRATION REPORT
Company: ABC Process Industries Ltd |  |  |  | Doc no: CAL TE 001 Rev: 0
Yellow cells: enter data. Grey cells: calculated automatically. Free template from instrumentationblog.in
Sensor details
Tag number |  | TE 2205 |  | Sensor type |  | Pt100 class A, 3 wire
Use range |  | 0 to 300 °C |  | Tolerance (± °C) |  | 0.5
Temperature source |  | Dry block calibrator |  | Reference thermometer |  | Reference PRT, certificate CC 24610
Calibration date |  | 2026-10-25 |  | Immersion depth (mm) |  | 150
Calibration results
Point | Set temperature (°C) | Reference (°C) | Sensor as found (°C) | As found error (°C) | As found result | Raw value (Ω or mV) | Stable (yes/no) | Remarks
1 | 0 | 0.02 | 0.1 | 0.08 | PASS | 100.04 | Yes
2 | 100 | 100.05 | 100.21 | 0.16 | PASS | 138.59 | Yes
3 | 200 | 199.96 | 200.3 | 0.34 | PASS | 175.98 | Yes
4 | 300 | 300.08 | 300.52 | 0.44 | PASS | 212.18 | Yes | Near limit
5
6
Maximum error (°C) | [6 empty] | 0.44
Overall result | [6 empty] | PASS
Wait for the block and both thermometers to stabilise at each point. Use an insert that fits the sensor closely and immerse it fully. Sensors cannot be adjusted; a failed sensor is replaced or corrected in the transmitter.
Performed by |  |  | Checked by |  |  | Approved by
Name: |  |  | Name: |  |  | Name:
Signature: |  |  | Signature: |  |  | Signature:
Date: |  |  | Date: |  |  | Date: