## How to Use
HOW TO USE: TEMPERATURE TRANSMITTER CALIBRATION REPORT
Colour legend
Yellow cells: type your data here.
Grey cells: formulas. They calculate automatically; do not overwrite them.
Steps
1. Enter the tag, sensor type and calibrated range (LRV and URV in °C) and the tolerance.
2. Disconnect the sensor and connect the calibrator in RTD (Ω) or thermocouple (mV) simulation mode.
3. For each test point simulate the Applied temperature, write the simulated Ω or mV in the yellow column and the transmitter output in As found output (mA).
4. Adjust (sensor trim or output trim) if any point fails and record the As left readings.
5. Reconnect the sensor and confirm a sensible process reading before closing the work order.
What the formulas do
• Applied temperature = LRV + (URV − LRV) × test point % ÷ 100.
• Expected output = 4 + 16 × test point % ÷ 100 mA.
• Error (% of span) = (measured mA − expected mA) ÷ 16 × 100; PASS when within the tolerance.
Sheets: "Template" is blank and ready to use. "Example" shows a completed form so you can see the expected format.
More free templates, calculators and tools: https://instrumentationblog.in/downloads/templates/

## Template
TEMPERATURE TRANSMITTER CALIBRATION REPORT
Company: Your Company Name / Logo |  |  |  |  | Doc no: Rev:
Yellow cells: enter data. Grey cells: calculated automatically. Free template from instrumentationblog.in
Instrument details
Tag number |  |  |  |  | Service
Make and model |  |  |  |  | Serial number
Sensor type |  |  |  |  | Engineering unit
Calibrated range LRV |  |  |  |  | Calibrated range URV
Tolerance (± % of span) |  |  |  |  | Damping (s)
Test equipment and conditions
Simulator / calibrator |  |  |  |  | Standard certificate no
Calibration date |  |  |  |  | Ambient temperature (°C)
Calibration results (sensor simulated at the transmitter input)
Test point (% of span) | Applied temperature | Simulated signal (Ω or mV) | Expected output (mA) | As found output (mA) | As found error (% span) | As found result | As left output (mA) | As left error (% span) | As left result
0 |  |  | 4
25 |  |  | 8
50 |  |  | 12
75 |  |  | 16
100 |  |  | 20
Maximum as found error (% span)
Maximum as left error (% span)
Overall result (as left if adjusted, else as found)
Pt100 resistance values from the IEC 60751 table (0, 50, 100, 150, 200 °C). For thermocouples enter the mV from IEC 60584 and set the calibrator to the correct cold junction mode.
Performed by |  |  | Checked by |  |  | Approved by
Name: |  |  | Name: |  |  | Name:
Signature: |  |  | Signature: |  |  | Signature:
Date: |  |  | Date: |  |  | Date:

## Example
TEMPERATURE TRANSMITTER CALIBRATION REPORT
Company: ABC Process Industries Ltd |  |  |  |  | Doc no: CAL TT 2101 01 Rev: 0
Yellow cells: enter data. Grey cells: calculated automatically. Free template from instrumentationblog.in
Instrument details
Tag number |  | TT 2101 |  |  | Service |  | Reactor R 201 jacket outlet
Make and model |  | Yokogawa YTA110 |  |  | Serial number |  | S5N812443
Sensor type |  | Pt100, 3 wire (IEC 60751) |  |  | Engineering unit |  | °C
Calibrated range LRV |  | 0 |  |  | Calibrated range URV |  | 200
Tolerance (± % of span) |  | 0.25 |  |  | Damping (s) |  | 2
Test equipment and conditions
Simulator / calibrator |  | Fluke 726 |  |  | Standard certificate no |  | NABL/CC 24203
Calibration date |  | 2026-10-12 |  |  | Ambient temperature (°C) |  | 27
Calibration results (sensor simulated at the transmitter input)
Test point (% of span) | Applied temperature | Simulated signal (Ω or mV) | Expected output (mA) | As found output (mA) | As found error (% span) | As found result | As left output (mA) | As left error (% span) | As left result
0 | 0 | 100 | 4 | 4.006 | 0.0375 | PASS | 4.001 | 0.00625 | PASS
25 | 50 | 119.4 | 8 | 8.012 | 0.075 | PASS | 8.002 | 0.0125 | PASS
50 | 100 | 138.51 | 12 | 12.02 | 0.125 | PASS | 12.001 | 0.00625 | PASS
75 | 150 | 157.33 | 16 | 16.026 | 0.1625 | PASS | 16 | 0 | PASS
100 | 200 | 175.86 | 20 | 20.03 | 0.1875 | PASS | 20.002 | 0.0125 | PASS
Maximum as found error (% span) | [7 empty] | 0.1875
Maximum as left error (% span) | [7 empty] | 0.0125
Overall result (as left if adjusted, else as found) | [7 empty] | PASS
Pt100 resistance values from the IEC 60751 table (0, 50, 100, 150, 200 °C). For thermocouples enter the mV from IEC 60584 and set the calibrator to the correct cold junction mode.
Performed by |  |  | Checked by |  |  | Approved by
Name: |  |  | Name: |  |  | Name:
Signature: |  |  | Signature: |  |  | Signature:
Date: |  |  | Date: |  |  | Date: